Automated Orbital Wall Fracture Detection and Severity Classification of Ocular Injuries Using Deep Learning

Contributors:

Farhad Salari1, Zahra Meskar2, Farhad Fallah2, Maryam Fazeli2, Reza Samiee1, Hossein Arabalibeik3, Seyed Mohsen Rafizadeh1*

Departments and institutions:

1-Eye Research Center, Farabi Eye Hospital, Tehran University of Medical Sciences, Tehran, Iran

2- Department of Electrical Engineering, Sharif University of Technology, Tehran, Iran

3- Research Center for Biomedical Technologies and Robotics (RCBTR), Tehran University of Medical Sciences, Tehran, Iran

* Corresponding author:

Seyed Mohsen Rafizadeh

E-mail: mohsen_raf1354@yahoo.com

Abstract

Background: Orbital fractures are frequently encountered in maxillofacial trauma and can be associated with severe globe injuries (SGI). Accurate and timely diagnosis is critical to prevent long-term complications, yet subtle fractures and soft tissue injuries may be overlooked. Recent advances in artificial intelligence offer promising tools to support clinical decision-making in acute settings.

Methods: We collected CT images from 250 patients with orbital wall fractures treated at a single center between May 2020 and May 2023. From CT scans of 250 included subjects, a dataset of 18,000 annotated orbital CT scan slices (250 subjects) was used, with segmentation preprocessing based on a U-Net framework. Two supervised convolutional neural network (CNN) classifiers were trained on axial and coronal CT images of the orbit. The first model (FxDetectCNN) was trained to detect the presence of orbital fractures, while the second model (SGIDetectCNN) classifies the severity of injury into mild or severe categories. The models’ performance was compared with annotations by two experienced ophthalmologists.

Results: The fracture detection model achieved an overall accuracy of 81%, with 86% sensitivity and 76% specificity. The SGI estimation model achieved 64% accuracy, with a sensitivity of 80% and a specificity of 48%, compared with expert annotation, which reached a maximum accuracy of 59.5% for SGI estimation.

Conclusion: Our study demonstrates the feasibility of CNN models in detecting fractures and classifying the severity of ocular injuries from CT scans. These tools can assist emergency physicians in triaging patients and may serve as valuable decision-support systems in emergency settings.

Keywords: Orbital fracture, Artificial intelligence, CT imaging, Deep learning, Ocular injury, Globe injury

# Introduction

Traffic collisions and sports-related injuries have led to a growing incidence of orbital fractures. Orbital fractures are one of the most frequently encountered facial fractures in maxillofacial surgery, accounting for up to 90% of such cases and approximately 3% of emergency room consultations 1,2.  Accurate diagnosis of orbital fractures relies on orbital tomography (CT) 3,4. While larger fractures are typically easy to identify using CT, more subtle fractures may be missed by non-experts. Moreover, orbital fractures may result in severe globe injuries (SGI), which can be sight-threatening and require urgent ophthalmic intervention 5.

Many studies have investigated clinical scoring systems that combine clinical and demographic data with radiologic findings to enhance the detection and assessment of ocular injuries in patients with orbital wall fractures 5-13; however, none of these approaches have gained widespread use due to the challenges of applying them in real-world clinical settings. Recent advancements in deep learning methods have resulted in the development of models that can improve the detection of orbital blowout fractures 14,15. For instance, Li et al. utilized convolutional neural networks (CNNs) combined with an XGBoost model to detect orbital blowout fractures, reporting a single-image accuracy of 92% and an area under the curve (AUC) of 0.9574 16. Umapathy et al. focused on volumetric analysis by using a modified U-Net model to quantify globe volumes, achieving a high F1 score of 0.95 17. Also, Bao et al. employed DenseNet-169 along with U-Net for the identification and segmentation of orbital fractures 18. One of the limitations of these previous works is the lack of models trained and validated on complex, multi-wall orbital fractures.

Despite advances demonstrating the value of artificial intelligence (AI) in orbital trauma diagnostics, AI-based prediction of severe globe injuries (SGI) in orbital fractures remains unexplored. Timely access to ophthalmologists in emergency settings remains limited, hindering the diagnosis and management of ocular injuries associated with orbital fractures. Severe ocular injuries are at risk of being overlooked by non-ophthalmologists, as specialized examination instruments, such as slit lamps and indirect ophthalmoscopes, are not always readily available in emergency departments and require specific training to use. Failure to promptly recognize ocular injuries may lead to serious long-term complications. To bridge this care gap, we introduce the development of two separate AI models: one for fracture detection and the second for predicting SGI in subjects with orbital wall fractures. This model aims to improve triage by non-ophthalmologists while serving as a decision-support system for ophthalmologists, enabling more precise and timely diagnosis 19.

# Methods

This retrospective cohort study was conducted in accordance with the tenets of the Helsinki Declaration and was approved by the institutional ethics committee of Tehran University of Medical Sciences (IR.TUMS.FARABIH.REC.1401.026).

Dataset and annotations

Imaging from subjects diagnosed with orbital fractures at Farabi Eye Hospital between May 2020 and May 2023 were included in this study. Patients were eligible for inclusion if the interval from trauma to their first ophthalmologic examination was less than one week and if complete ophthalmologic evaluations, along with high-resolution fine-cut orbital CT imaging (2 mm slice thickness and 1 mm slice increment ) without contrast, were available. Subjects were excluded if visual acuity data were missing, whether due to patient non-cooperation (e.g., young age or inability to complete the Snellen chart) or incomplete documentation. Additionally, patients with a history of ocular surgery (e.g., corneoscleral repair, canthotomy, or cantholysis) were excluded.

Eyes with orbital fractures were labeled based on medical records documented by an experienced oculoplastic surgeon (SMR). The severity of ocular injury was determined from clinical records using a modification of Rossin's classification 6 (Supplementary Table 1)20. In this classification, ocular injuries associated with orbital fractures are categorized into four grades according to the urgency of ophthalmologic intervention. Injuries classified as grades 2 and 3 were considered severe globe injuries (SGI).

The dataset was randomly split into a training set (200 cases) and an internal validation set (50 cases). To establish a baseline for evaluating model performance against expert clinical judgment, two expert ophthalmologists (SMR and FS), blinded to the clinical information,  independently annotated the validation dataset using CT imaging alone. Both experts reviewed 50 randomly selected samples and had access to all available imaging planes, including sagittal, coronal, and axial views.

Preprocessing

The dataset used in this study consists of high-resolution orbital CT images with corresponding annotations indicating whether a fracture is present and, for fractured cases, the severity (mild or severe). 18000 orbital CT slices were obtained from 250 included patients diagnosed with orbital fractures. All images were preprocessed to ensure uniform input sizes and consistent normalization for effective model training.

We extracted the axial dataset from DICOM files and standardized it by selecting 40 slices per case, since the number of slices varied across patients. To focus on the region of interest (ROI), specifically the eye globe and its surrounding orbital bones, we used the proposed U-Net network [1] to segment the mask. For each case, we began by defining an all-zero tensor. Next, we passed each axial slice through the model, added the output mask to the zero tensor, and repeated these steps for all slices. Finally, we applied a nonzero filter to identify the largest connected area, which was assumed to contain the eye globe. The processing area was then refined by applying dilation operations to the extracted mask. To enhance focus on relevant details, we applied a Hounsfield Unit (HU) window of -50 to 50 to the output. Finally, we extracted both eye globes, concatenated them into a single image, and normalized the resulting image to prepare it for further processing (Figure 1).  At the next stage, each patient’s CT volume was processed to create two eye-specific samples (right and left eyes). Rather than splitting the image by fixed pixel ranges, we isolated the orbital region by centering a crop on the midline of each axial or coronal slice. Specifically, the image was cropped around the middle and then symmetrically padded by 128 pixels on both the left and right sides to preserve the central field of view and ensure a consistent input size focused on the mid-orbital anatomy. This standardization step retained the mid portion of the scan where the orbits are best visualized while minimizing irrelevant peripheral content. Horizontal flipping was applied as needed to maintain anatomical consistency between left- and right-eye samples.

For patients with unilateral orbital fractures, each eye was treated as an independent sample, with the affected eye labeled as fracture (1) and the unaffected eye labeled as normal (0). Both samples were included in the dataset, using the contralateral eye as a within-patient control. For patients with bilateral orbital fractures, both eyes were included as separate positive samples, each labeled as fracture (1). Labels for right and left eyes were extracted from the ground truth annotations and concatenated to form the final eye-level label set. This strategy increased the number of training instances while preserving the true fracture status of each eye and treating each orbit as an independent anatomical unit for classification.

Data augmentation techniques, including random rotations, flips, and zooms, were applied to increase the diversity of the training set and improve model robustness. Here, we aim to conduct two primary tasks: binary classification to detect the presence of a fracture (fracture vs. non-fracture) and binary classification to determine severity (severe vs. non-severe).

Model Architectures

We developed two convolutional neural network (CNN) architectures using 18,000 cropped CT images, focusing on the orbital region from the 250 included cases. The first, FxDetectCNN, integrates axial and coronal CT views via a dual-branch architecture to detect orbital fractures. The second, SGIDetectCNN, is a single-branch CNN that processes axial slices to classify fracture severity as severe or non-severe. This design allows the detection model to benefit from multi-view structural information, while the severity model emphasizes detailed axial features relevant to soft-tissue and optic nerve involvement.

The FxCNN_Model is designed to classify orbital CT images into fracture and non-fracture categories. It consists of two main branches, one processing the axial slice and the other the coronal slice. Both branches are equipped with convolutional layers, batch normalization, and max-pooling operations to extract spatial features from the input images. The two branches are then concatenated to create a joint representation, which is passed through a fully connected layer for final classification. The model structure is displayed in Figure 2. The SGIDetectCNN is intended to classify fractures as non-severe or severe. This model uses a similar structure to the fracture classification model (Figure 3).

The models were trained with a 75-25 train-test split, using the training set for optimization and the validation set for evaluation. Training employed Binary Cross Entropy Loss and Adam optimizer with an initial learning rate of 10⁻³, reduced by a factor of 10 every 5 epochs if validation loss did not improve. A batch size of 8 was chosen, with a maximum of 50 epochs and early stopping if validation loss did not improve for 10 consecutive epochs. L1 and L2 regularization, and dropout layers were applied to prevent overfitting. Model performance was tracked after each epoch, with the best model based on validation accuracy saved.

All experiments were conducted on a machine equipped with an NVIDIA GTX 1660 GPU, and total training time per model ranged from 5 to 10 minutes, depending on the number of epochs and model complexity.

Statistical Analysis

All statistical analyses were performed using the open-source software PyTorch (version 2.3.1) and Python 3.12.4. The data were presented as averages with standard deviations for continuous variables and as counts for categorical variables. Classification accuracy was evaluated using sensitivity, specificity, and F1 score (harmonic mean of recall and precision).

# Results

250 subjects met the eligibility criteria, comprising 205 males (81.9%) and 45 females (18.1%), with a mean age of 32.9 ± 14.42 years (range: 5–79 years). Details of demographic and radiologic features of included subjects are summarized in Supplementary Table 2.

The dataset comprised 200 training and 50 internal validation volumetric orbital CT scans, each annotated with the affected eye's laterality. Each CT scan consisted of an axial volumetric series centered on the orbital region. Laterality information was used to standardize anatomical orientation during preprocessing, as described in the Methods section.

Fracture detection

The FxDetectCNN achieved an overall accuracy of 81% in detecting orbital bone fractures, with a sensitivity of 86% and specificity of 76%. In addition, the F1 score was 0.82, indicating substantial agreement between model predictions and ground-truth annotations. These results demonstrate that the model performs particularly well at detecting severe or positive injury cases, a critical capability in clinical applications where the cost of false negatives can lead to significant diagnostic and therapeutic delays and critical diagnostic oversights. Figure 4 presents the confusion matrix for orbital fracture detection, and Supplementary Figure 1 shows the training and validation performance across 20 epochs.  For orbital fracture detection, the FxDetectCNN achieved an area under the ROC curve (AUC) of 0.83.

SGI Estimation

A separate network, SGIDetectCNN, was trained to classify orbital injury severity as severe or non-severe. The model achieved an overall accuracy of 64%, with a sensitivity of 80%, a specificity of 48%, and an F1 score of 0.69. This indicates a strong ability to detect severe cases, although the model showed limitations in differentiating non-severe injuries. We compared the SGIDetectCNN model's performance with other classification models, demonstrating its superiority with an F1 score of 0.69. Supplementary Figure 2 shows the training and validation performance across 20 epochs. For severe globe injury (SGI) estimation, the SGIDetectCNN achieved an AUC of 0.75.

Comparison with experts’ performance

The first expert (FS) achieved 94.0% accuracy in binary injury detection and 56.5% in injury severity estimation. The second expert (SMR) demonstrated comparable accuracy for fracture detection (94.0%) and slightly higher accuracy for severity estimation (59.5%). Figures 4 and 5 illustrate the expert-annotated outcomes for injury severity estimation.

To clarify the model’s decision-making process, we applied Gradient-weighted Class Activation Mapping (Grad-CAM) to generate visual explanations of the predictions. In fracture cases, the focus was primarily on the bony structures of the orbit (Figure 6). In the severity estimation model, Grad-CAM revealed concentrated attention in clinically relevant areas when severe injuries were present (Figure 7). By contrast, in mild cases, the attention maps were more diffuse and extended into peripheral regions, which are less indicative of severity (Figure 7). This pattern suggests that while the model reliably localizes diagnostic features in severe injuries, its interpretability and confidence decrease when distinguishing more subtle presentations.

Discussion

Fine-cut orbital CT scans are the gold standard for diagnosing orbital wall fractures and for planning their management 21. Although a comprehensive ophthalmic examination is the most reliable method for diagnosing severe ocular injuries associated with orbital wall fractures, factors such as decreased level of alertness or coexisting conditions, including severe periorbital edema, may limit emergency physicians' ability to accurately assess and manage ocular injuries 22. Additionally, the intraconal and extraconal fatty pads around the globe provide excellent contrast, revealing details of eye structures. Recent studies also provided clues that certain aspects of CT imaging can be used to predict SGIs, including the location of the fracture 6 and the muscle involvement 20.

The present study demonstrates the potential of an interpretable CNN framework for the dual tasks of orbital fracture detection and injury severity estimation from CT scans. Our model achieved 81% accuracy for fracture detection and 64% for severity estimation. The implementation of Grad-CAM provided crucial model interpretability, confirming that its diagnostic focus aligns with clinically relevant anatomical structures. The severity estimation model demonstrated similar or better accuracy than the diagnostic performance of human experts (56.5%–59.5%), suggesting that such AI systems can be used to reduce inter-observer variability and enhance diagnostic reliability in orbital trauma assessment.

Most applications of artificial intelligence in orbital fractures have focused on detecting and segmenting blowout fractures. Kim et al. (2025) achieved an accuracy of 99.47% for blowout fracture detection 23, while studies by Bao et al. (2023) and Li et al. (2020) also reported accuracies exceeding 92% using various deep learning architectures 16,18.  The work by Li et al. (2024), which utilized a dual-network system (YOLOv7 for detection and an enhanced UNet for segmentation), achieved a detection F1-score of approximately 76.8% and a high segmentation Dice coefficient of 90.2% 24. The hierarchical system developed by Oku et al. (2025) achieved excellent performance in classifying fracture morphology, with patient-level sensitivity of 94.7% for detection and 100% for differentiating clinically urgent trap-door fractures 25. While these models demonstrated excellent performance, their scope was often confined to specific fracture subtypes. A significant contribution of this research is its departure from the current focus on isolated blowout fractures in the existing literature. By developing a model trained on an inclusive dataset encompassing all fracture morphologies, including complex multi-wall patterns, our framework possesses enhanced ecological validity. The FxDetectCNN model showed an acceptable performance for the detection of fractures in a dataset of 250 cases with mixed orbital wall fractures; however, further model adjustment on a larger dataset is required to improve the model’s accuracy.

To our knowledge, this study is the first to apply deep learning to improve SGI prediction in patients with orbital fractures. Estimating injury severity using only orbital CT scans has inherent limitations; for example, mild anterior segment hemorrhage or small full-thickness corneal lacerations cannot be detected with imaging alone. However, the widespread availability of CT in emergency departments and the difficulty in performing full ocular examinations by non-ophthalmologists in acute settings justifies efforts to enhance its utility by incorporating deep learning techniques. The SGIDetectCNN model achieved 64% accuracy on this complex task, surpassing the rates of expert clinicians. Although the specificity for severity estimation was moderate (48%), the model’s high sensitivity (80%) underscores its potential as a robust screening tool for the early identification of severe cases where timely intervention is paramount. Future studies are suggested to evaluate model performance in a prospective cohort to demonstrate its potential to improve clinical outcomes.

This study has several limitations. The model was developed using a modest sample size of 250 cases and 18000 CT slices from a single institution, which may limit the generalizability of our findings.  We suggest that further studies on model performance be conducted in a multicenter setting to externally validate this work. The observed variability in expert annotations for severity grading highlights the intrinsic subjectivity of the SGI diagnostic task. Finally, the model input in this paper was only orbital CT scans, which have inherent limitations in the diagnosis of subtle ocular injuries. Future work could focus on building models that integrate optical imaging, such as fundus photography and optical coherence tomography, with CT scans.

# This study introduces a novel deep learning framework that can detect orbital fractures and estimate injury severity. These models achieved clinically relevant performance, with the severity estimation model performing at least as well as expert ophthalmologists. This work is a first step toward applying AI as a decision-support tool in the diagnosis and management of orbital wall fractures in emergency settings where specialized expertise may not be immediately available.

References:

1	Scolozzi, P. Reflections on a patient-centered approach to treatment of blow-out fractures: Why the wisdom of the pastmust guide our decision-making. Journal of Plastic, Reconstructive & Aesthetic Surgery 75, 2268-2276, doi:https://doi.org/10.1016/j.bjps.2022.04.034 (2022).

2	Bord, S. P. & Linden, J. Trauma to the Globe and Orbit. Emergency Medicine Clinics of North America 26, 97-123, doi:https://doi.org/10.1016/j.emc.2007.11.006 (2008).

3	Rhee, S. H., Kim, T. S., Song, J. M., Shin, S. H. & Lee, J. Y. Effectiveness of Computed Tomography for Blow-out Fracture. Maxillofac Plast Reconstr Surg 36, 273-279, doi:10.14402/jkamprs.2014.36.6.273 (2014).

4	Santos, D., Oliveira, J., Vannier, M. & Cavalcanti, M. Computed tomography imaging strategies and perspectives in orbital fractures. Journal of applied oral science : revista FOB 15, 135-139, doi:10.1590/S1678-77572007000200012 (2007).

5	Zhong, E. et al. Orbital Fractures and Risk Factors for Ocular Injury. Clinical Ophthalmology, 4153-4161 (2022).

6	Rossin, E. J. et al. Factors Associated With Increased Risk of Serious Ocular Injury in the Setting of Orbital Fracture. JAMA Ophthalmol 139, 77-83, doi:10.1001/jamaophthalmol.2020.5108 (2021).

7	Andrews, B. T. et al. Orbit fractures: identifying patient factors indicating high risk for ocular and periocular injury. The Laryngoscope 126, S5-S11 (2016).

8	Chow, J., Parthasarathi, K., Mehanna, P. & Whist, E. Primary assessment of the patient with orbital fractures should include pupillary response and visual acuity changes to detect occult major ocular injuries. Journal of oral and maxillofacial surgery 76, 2370-2375 (2018).

9	Ho, T. Q., Jupiter, D., Tsai, J. H. & Czerwinski, M. The incidence of ocular injuries in isolated orbital fractures. Annals of plastic surgery 78, 59-61 (2017).

10	Layton, C. J. Factors associated with significant ocular injury in conservatively treated orbital fractures. Journal of ophthalmology 2014 (2014).

11	Richani, K. et al. Screening criteria for detecting severe ocular injuries in the setting of orbital fractures. Ophthalmic Plastic & Reconstructive Surgery 35, 609-614 (2019).

12	Dutton, G. et al. Ophthalmic consequences of mid-facial trauma. Eye 6, 86-89 (1992).

13	Lange, C., Feltgen, N., Junker, B., Schulze-Bonsel, K. & Bach, M. Resolving the clinical acuity categories “hand motion” and “counting fingers” using the Freiburg Visual Acuity Test (FrACT). Graefe's Archive for Clinical and Experimental Ophthalmology 247, 137-142 (2009).

14	Niño-Sandoval, T. C., Landim, F. S. & Vasconcelos, B. C. Deep Learning for Orbital Fracture Detection and Reconstruction: A Systematic Review on Diagnostic Accuracy and Surgical Planning. Journal of Cranio-Maxillofacial Surgery (2025).

15	Ekman, T., Barakat, A. & Heiberg, E. Generalizable deep learning framework for 3D medical image segmentation using limited training data. 3D Printing in Medicine 11, 1-18 (2025).

16	Li, L. et al. Deep Convolutional Neural Networks for Automatic Detection of Orbital Blowout Fractures. J Craniofac Surg 31, 400-403, doi:10.1097/scs.0000000000006069 (2020).

17	Umapathy, L. et al. Fully Automated Segmentation of Globes for Volume Quantification in CT Images of Orbits using Deep Learning. AJNR Am J Neuroradiol 41, 1061-1069, doi:10.3174/ajnr.A6538 (2020).

18	Bao, X. L. et al. Automatic Identification and Segmentation of Orbital Blowout Fractures Based on Artificial Intelligence. Transl Vis Sci Technol 12, 7, doi:10.1167/tvst.12.4.7 (2023).

19	Wawer Matos, P. A. et al. Artificial Intelligence in Ophthalmology - Status Quo and Future Perspectives. Semin Ophthalmol 38, 226-237, doi:10.1080/08820538.2022.2139625 (2023).

20	Salari, F. et al. Prediction of substantial closed-globe injuries in orbital wall fractures. Int Ophthalmol 44, 219, doi:10.1007/s10792-024-03113-w (2024).

21	Boyette, J. R., Pemberton, J. D. & Bonilla-Velez, J. Management of orbital fractures: challenges and solutions. Clin Ophthalmol 9, 2127-2137, doi:10.2147/opth.S80463 (2015).

22	Lee, H. J., Jilani, M., Frohman, L. & Baker, S. CT of orbital trauma. Emerg Radiol 10, 168-172, doi:10.1007/s10140-003-0282-7 (2004).

23	Kim, S., Koh, H. K., Lee, H. & Shin, H. J. Deep-Learning Method for the Diagnosis and Classification of Orbital Blowout Fracture Based on Computed Tomography. Journal of Oral and Maxillofacial Surgery, doi:10.1016/j.joms.2025.04.010.

24	Li, Y.-L., Li, Y.-H., Wei, M.-Y. & Li, G.-Y. Automatic detection of optic canal fractures and recognition and segmentation of anatomical structures in the orbital apex based on artificial intelligence. Frontiers in Cell and Developmental Biology Volume 13 - 2025, doi:10.3389/fcell.2025.1609028 (2025).

25	Oku, H. et al. Hierarchical deep learning system for orbital fracture detection and trap-door classification on CT images. Computers in Biology and Medicine 196, 110732, doi:https://doi.org/10.1016/j.compbiomed.2025.110732 (2025).

Funding:

We received no funding for the completion of this study.

# Competing intrests:

Authors declare no conflict of interest.

# Code and Data Availability

The full implementation and training scripts are publicly available at: https://github.com/Nadiam75/orbital-injury-detection/. The datasets generated and/or analyzed during the current study are not publicly available due to patient privacy and ethical restrictions, but are available from the corresponding author upon reasonable request.

Author contributions:

FS: conception, data extraction, validation, data analysis, writing - review & editing; ZM.: writing, generated the figures; FF: data analysis, model architecture design; MF: data analysis, model architecture design; RS: data extraction, validation, writing;  HA: data curation, validation, supervision, resources; SMR: data annotation, validation, supervision, resources; All authors reviewed and edited the manuscript

Figure legends:

Figure 1. Preprocessing steps for axial CT images.

Figure 2. Model structure for fracture detection (FxDetectCNN)

Figure 3. Model structure for detection of severe ocular injuries ( SGIDetectCNN).

Figure 4. Confusion matrix for orbital fracture detection (left) and injury severity estimation (right).

Figure 5. Expert annotations by (a): First expert: 56.5% (b): Second expert: 59.5% for the task of injury severity estimation.

Figure 6. Grad-CAM visualizations highlighting model attention for fracture detection in two cases. (a) Case 1 without fracture and (b) Case 2 with medial and inferior orbital wall fractures.

Figure 7. Grad-CAM Heatmaps for model predictions for severity estimation in two cases. (a) Case 3 with a mild ocular injury and (b) Case 4 with a severe ocular injury.